## Slide 1
GOES-13 Imager
Corrupt Space Look causing bad 4 um data?
(affecting at least fog product)
Path forward….
Timothy J. Schmit
(NOAA/NESDIS/STAR)
Mathew M. Gunshor, James P. Nelson III and Anthony J. Schreiner
(CIMSS)
FangFang Yu (ERT, Inc.)
Michael Sporer (NWS)
17-May-2013

## Slide 2: Outline
Background
Examples
  August 2012
  April 2013
Way forward options
Requirements

## Slide 3: Background
Historically, the GOES-8/12 series did not take images during “KOZ”
  Could be up to 3 hours
GOES-13/15 allows images to now be taken during this time period
Recent issue seems to be corrupted space look and not directly related to the stray light correction
Space looks are scheduled to be on the ‘clean’ side of the earth (away from the sun, etc.)

## Slide 4: Comment from NWS
From: Patrick Ayd <patrick.ayd@noaa.gov> Date: Sun, Aug 19, 2012 at 1:24 AM Subject: INFO: Bad Fog Channel Images To: Joshua Scheck <Joshua.Scheck@noaa.gov>
Josh,
I have noticed more and more during the night that bad 3.9 mu satellite images are corrupting the 11-3.9 mu fog channel, reversing the way clouds should appear (i.e all clouds both low and high are colored as very dark).
I have attached two images from 545 UTC on 8-19-12:
fog_sat.png: 11-3.9 mu fog channel with cig/vis plot
3_9_mu.png: 3.9 mu channel
Thanks,
-Patrick
--
Patrick Ayd
General Forecaster
NOAA/NWS WFO Bismarck, ND
Office Phone: (701) 223-4582
2012:

## Slide 5
<-Note seam
11 – 4 um in counts
11 um
6.5 um
4 um
<-Note seam
AWIPS as seen at CIMSS

## Slide 6
AWIPS as seen at CIMSS
Bad data

## Slide 7
AWIPS as seen at CIMSS

## Slide 8
AWIPS as seen at CIMSS
Bad data

## Slide 9
AWIPS as seen at CIMSS

## Slide 10
From the NWS (Blacksburg)
2013:

## Slide 11
From the NWS (Blacksburg)

## Slide 12
From the NWS (Blacksburg)
Notes: Should be CIMSS. 

## Slide 13
Visible band:

## Slide 15
Note stray light evident in visible band

## Slide 16
Note stray light evident in visible band

## Slide 18
Imager, band 4, longwave window

## Slide 19
Imager, band 4, longwave window

## Slide 20
Imager, band 4, longwave window

## Slide 21
Imager, band 4, longwave window

## Slide 22
Imager, band 4, longwave window

## Slide 23
Imager, band 4, shortwave window

## Slide 24
Imager, band 4, shortwave window

## Slide 25
Imager, band 4, shortwave window

## Slide 26
Imager, band 4, shortwave window
Brightness temperatures corrupted at this time

## Slide 27
Imager, band 4, shortwave window

## Slide 28
Time difference (band 2)
Artificially colder by ~50K

## Slide 29
Time difference (band 4)
Notes: Of course these shouldn’t be the same, given 45 minutes between images.

## Slide 30
From Fangfang Yu
“Out-of-family” Space looks at 05 UTC

## Slide 37: Way Forward Possible Options
Do nothing
  No good, since products are affected and users have noticed.
Modify calibration procedure
  Filter space look or use previous good space look
  Clamp space look only if post-clamp count is higher than pre-clamp count (X. Wu)
  Scan elsewhere for good space look data, etc.
    Scan to opposite hemisphere during these times
Have product producers check GVAR stray light flags or Notice to Users
  Only works on case-by-case basis
  Each PAL would need to be involved
  NWS/AWIPS does not have the GVAR flags
  Still would produce non-spec compliant 4 um images
    Spec <1K from truth
Cancel/Shift frames (with ‘corrupt’ space views):
  when the sun is right above/below the poles
  Less than one per day (2 seasons)

## Slide 38: Output from line prefix
STRAYLIGHT 213
(STRAYLIGHT.PGM) *START execution on 2013129 (Thursday May 9, 2013) at 181216 (HHMMSS) UTC
(STRAYLIGHT.PGM) Input program commandline (in 80-byte segments):
STRAYLIGHT 213
(STRAYLIGHT.FOR) OPEN/READ GOES areafile: /home/tims/mcidas/data/AREA0213
(STRAYLIGHT.FOR) Area 213 directory information (Image type, SSEC sensor source #, image ccyyddd, image hhmmss):
4, 180, 2013111, 50200
(STRAYLIGHT.FOR) Check each image line for Stray Light Correction (SLC) information.
(STRAYLIGHT.FOR) Input area AREA0213 lines contain prefixes -- size= 488 bytes= 122 words
(STRAYLIGHT.FOR) Size of prefix DOC (4-byte words)= 119, CAL= 0, BND= 2
(STRAYLIGHT.FOR) Read input area AREA0213 line prefixes now.
(STRAYLIGHT.FOR) All 800 AREA0213 line prefixes now read.
(STRAYLIGHT.FOR) SLCE (# lines w/SLC enabled) = 784, SLCD (# lines w/SLC disabled) = 16
(STRAYLIGHT.FOR) SLCP (# lines w/SLC performed) = 784, SLCNP (# lines w/SLC not performed) = 16
(STRAYLIGHT.FOR) SLCP_FIRST (1st line w/SLC performed) = 17, SLCP_LAST (last line w/SLC performed) = 800
(STRAYLIGHT.FOR) SUNE (# lines w/sun on e/side of earth) = 784, SUNW (# lines w/sun on w/side of earth) = 16
(STRAYLIGHT.FOR) SUNN (# lines w/sun on n/side of earth) = 784, SUNS (# lines w/sun on s/side of earth) = 16
(STRAYLIGHT.FOR) 784 nonzero values were found for the parameter: min_sun_angle_range_of_earth_pixels_in_a_scan
(STRAYLIGHT.FOR) Smallest value = 3 found at line # 17
(STRAYLIGHT.FOR) Largest value = 8 found at line # 669

## Slide 39: System Absolute Accuracy
Notes: From Databook