## Slide 1
GOES-13/14/15 imagers and sounders are capable of scanning the sun without health and safety issues. However, NOAA has to cancel or replace all imaging within 6 degrees due to intolerable sun intrusion.
Annual loss: Equivalent to 6 days of imaging.
The sun intrusion is more detectable on shorter wavelength IR channels (especially Channel 2) of the imager with the effect increasing as the scan mirror line-of-sight (LOS) gets closer to the sun.
NOAA and ITT Industries have characterized the effect of the sun intrusion and developed a correction algorithm to claim >95% of lost images.
GOES Imager Stray Light correction

## Slide 2
The ground system algorithm consists of:
  Subtracting the known amount of additional stray light from the signal outside 6°.
  Using the longer wavelength signal in combination with 3.9-um signal to estimate a truer 3.9-um signal in areas where the sun is within 6° of the pixel and the stray light effect is overwhelming (hybrid).
GOES Imager Stray Light correction

## Slide 3: GOES Imager Stray Light correction
GOES Imager Operational as of February 22, 2012 at 04 UTC.
Currently, GOES-13 Imager only is being corrected for stray light.
Mostly affects imager band 2 (4 micrometer), but all bands can be affected.
Allows extra images to be scanned around satellite midnight.
Corrects most of the ‘extra’ or stray radiation in the earth scenes
  Algorithm from ITT
Stray Light correction status bits sent out over GVAR and are now part of the McIDAS area line prefix
  Enabled, performed relative position of the sun, etc.
GOES-15 Imager stray light correction is planned for the fall of 2012
The GOES-R ABI can’t scan as close to the sun as the GOES-13/14/15

## Slide 9
GOES-13 Imager Band 2
Stray Light Correction
Original
Corrected
Later image
Later image
Notes: The 4 um (band 2 GOES imager) data can be over 50K too hot due to stray light. After being corrected, the image is more similar to subsequent (unaffected) images.
Data ingested by the SSEC data center.

## Slide 10
Band 2, Original – too hot by several K!
Band 2, Original – too hot by 50K!

## Slide 11
Band 2, Corrected
Notes: Ignore missing lines, due to ingest.

## Slide 12
Band 2, later image
Notes: Ignore missing lines, due to ingest.

## Slide 13
Band 2, difference image (org-eclipse). Smaller range.
Band 2 Difference
N=14342
Mean= 4.0K
Notes: Note that even far from the sun, there’s still a 4K difference.

## Slide 14
Band 2, difference image (corrected-eclipse). Smaller range.
Band 2 Difference
N=14342
Mean= 0.6K

## Slide 15
Cloud-top Pressure (color-range). Original – too “cloudy” over some regions, but misses other clouds!
Clear Low cloud High Cloud
Notes: Derived Product of cloud not correct with data with stray light.

## Slide 16
Cloud-top Pressure (color-range). Corrected. In general, corrected both regions.
Clear Low cloud High Cloud
Notes: …outside of 6 degrees.

## Slide 17
Cloud-top Pressure (color-range). Later image.
Clear Low cloud High Cloud
Notes: Derived Product of cloud much improved after stray light correction.

## Slide 18
Notes: Data from SSEC Data Center

## Slide 19
Notes: Data from SSEC Data Center

## Slide 20
Notes: Data from SSEC Data Center

## Slide 21
Notes: Data from SSEC Data Center

## Slide 22
Notes: Data from SSEC Data Center

## Slide 25
Notes: Data from SSEC Data Center

## Slide 26: GOES Imager Stray Light correction
Extra full disk image with the Stray light correction, compared to last year.
March 3, 2011
March 3, 2012
Notes: Of course other times are also added, not just this one time.

## Slide 27: GOES-11 (West) Imager Coverage
JD=077; March 18, 2011 (08:00 to 10:11 UTC); GOES-11 Imager

## Slide 28: GOES-15 (West) Imager Coverage
JD=077; March 17, 2012 (08:00 to 10:11 UTC); GOES-15 Imager

## Slide 29: SLC info found the GVAR stream
An example of the type of output of Stray Light Correction information found the GVAR stream, in the area line prefix for the GOES-13 Imager data is shown below. This provides a summary of how each image is modified depending on the time of year (i.e. relative declination of the sun and GOES platform).
Directory information (sensor #, image ccyyddd, image hhmmss): 180, 2012068, 54500
SLCE (# lines w/SLC enabled) = 2705
SLCD (# lines w/SLC disabled) = 0
SLCP (# lines w/SLC performed) = 2705
SLCNP (# lines w/SLC not performed) = 0
SLC_FRST (1st ln w/SLC performed) = 1
SLC_LST (last ln w/SLC performed) = 2705
SUNE (# lines w/sun on e/side earth) = 2705
SUNW (# lines w/sun on w/side earth) = 0
SUNN (# lines w/sun on n/side earth) = 0
SUNS (# lines w/sun on s/side earth) = 2705
2705 nonzero values were found for parameter: min_sun_angle
Smallest sun angle value = 5 found at line # 1419
Largest sun angle value = 15 found at line # 1

## Slide 30: Links
Links
http://cimss.ssec.wisc.edu/goes/blog/archives/9816
http://www.osd.noaa.gov/GVAR_Downloads/documents/SLC_GVAR_White_Paper_V1.doc
http://www.ssd.noaa.gov/PS/SATS/MESS/MSG0522058.01.txt

## Slide 31: GOES Imager Stray Light correction
http://cimss.ssec.wisc.edu/goes/blog/archives/9816
Notes: CIMSS blog web page about the SLC implementation.

## Slide 32: Back-up
Back-up

## Slide 33
Uncorrected

## Slide 34
Corrected

## Slide 35
Corrected - Sun within 1 degree

## Slide 36
Corrected - Sun within 2 degrees

## Slide 37
Corrected

## Slide 38
Uncorrected

## Slide 39
Uncorrected